# Supplementary material

Supplement to: The recording system's noise floor drives automated scoring of REM sleep without atonia: four pre-specified analyses of 1,354 open polysomnograms. Numbers are from the pipeline's reports; nights with at least 10 min of REM sleep; published cut-offs unless stated; FC = floor-corrected with B = 1.0 µV.

## S1. Participant flow (Supplementary Figure S1)

Supplementary Figure S1 · participant flow, six cohorts

Of 1,462 recordings, 1,354 nights were analyzed: 1,177 from the development cohorts and 177 from the external cohorts used only for the pre-specified test. HMC's three withdrawn recordings were removed by the database's providers for erroneous signal data; BOAS's 28 later nights were used only in a sensitivity analysis. Same-system comparisons use CAP's 21 RBD and 51 other patients recorded on its main system.

## S2. Exploratory analyses

None of the analyses in this section was fixed in advance.

Supplementary Table S1. Share flagged by the published cut-offs in each CAP group and the clinical cohorts, % [95% CI]. All 22 CAP RBD patients; native band.

| Method | CAP RBD (22) | CAP healthy (15) | NFLE (40) | PLM (10) | Insomnia (9) | Narcolepsy (5) | MGH (925) | HMC (147) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Montreal | 91 [72-97] | 13 [4-38] | 22 [12-38] | 40 [17-69] | 22 [6-55] | 60 [23-88] | 29 [27-32] | 9 [5-15] |
| SINBAR | 95 [78-99] | 60 [36-80] | 75 [60-86] | 70 [40-89] | 67 [35-88] | 100 [57-100] | 76 [73-79] | 25 [19-33] |
| RAI | 55 [35-73] | 25 [9-53] | 8 [3-20] | 10 [2-40] | 22 [6-55] | 20 [4-62] | 9 [7-11] | 7 [4-13] |

In a common 10-45 Hz band available to every recording, positivity was similar (Montreal MGH 29%, HMC 12%; SINBAR 84% and 29%). Discrimination between CAP RBD and same-system patients (AUC): Montreal tonic 0.844, phasic 0.846; SINBAR tonic 0.845, phasic 0.884, any 0.850; AASM tonic 0.685, phasic 0.852, phasic-epoch 0.839; RAI 0.842.

Supplementary Table S2. Agreement between methods (Cohen's kappa) and share of nights on which the three methods disagree.

| Group | Montreal-SINBAR | Montreal-RAI | SINBAR-RAI | Methods disagree, % of nights |
| --- | --- | --- | --- | --- |
| MGH (919) | 0.20 | 0.34 | 0.04 | 70 |
| HMC (147) | 0.36 | 0.73 | 0.29 | 22 |
| MGH with 1.5 µV of added noise (277) | - | 0.77 | - | 28 |
| CAP RBD (22) | 0.65 | 0.21 | 0.11 | 41 |

Supplementary Table S3. Clinical covariates in MGH. Logistic regression of a positive call (odds ratio [95% CI]) and share flagged by apnea-hypopnea index (AHI), with all REM epochs and with REM epochs containing a scored respiratory event or arousal excluded.

| Method | Per 10 years of age | Male sex | Per 10 AHI events/h |
| --- | --- | --- | --- |
| Montreal | 1.03 [0.93-1.14] | 1.16 [0.85-1.58] | 1.16 [1.05-1.28] |
| SINBAR | 0.89 [0.79-0.99] | 1.43 [1.04-1.97] | 1.19 [1.05-1.35] |
| RAI | 0.94 [0.80-1.11] | 0.81 [0.50-1.31] | 1.21 [1.05-1.39] |

| Share flagged, % | AHI below 5 (132) | 5-15 (276) | 15-30 (387) | 30 or more (130) |
| --- | --- | --- | --- | --- |
| Montreal, all REM | 24 | 27 | 31 | 35 |
| Montreal, event-free REM | 21 | 20 | 24 | 19 |
| SINBAR, all REM | 69 | 76 | 76 | 82 |
| SINBAR, event-free REM | 65 | 69 | 67 | 70 |
| RAI, all REM | 6 | 9 | 8 | 12 |
| RAI, event-free REM | 6 | 9 | 7 | 7 |

Excluding event epochs lowered overall positivity modestly (Montreal 28% to 22%, SINBAR 75% to 68%, RAI 7% to 7%) and removed the apnea gradient. By age (below 40, 40-60, above 60 years): Montreal 28, 29 and 31%; SINBAR 76, 78 and 74%; RAI 11, 8 and 9%.

Supplementary Table S4. Recordings re-made in software: share flagged (%) and paired changes.

| Copy | Quiet-REM level, µV | Montreal | SINBAR | RAI | Methods disagree, % |
| --- | --- | --- | --- | --- | --- |
| HMC as recorded | 1.52 | 9 | 25 | 7 | 22 |
| HMC, 1-µV steps | 1.54 | 9 | 24 | 7 | 21 |
| HMC, 1-µV steps and 200 Hz | 1.51 | 9 | 22 | 7 | 19 |
| HMC scaled by 0.39 | 0.59 | 8 | 25 | 3 | 23 |
| HMC scaled, 1-µV steps and 200 Hz | 0.63 | 8 | 20 | 3 | 18 |
| MGH as recorded | 0.59 | 29 | 76 | 9 | 70 |
| MGH (first 300) as recorded | 0.56 | 34 | 83 | 12 | 70 |
| MGH (first 300) with 1.5 µV of noise | 1.37 | 10 | 36 | 10 | 28 |

Paired, same nights (difference [95% CI]; nights changed; exact McNemar p): HMC in MGH's format versus HMC, SINBAR -2.7 [-5.4 to -0.1], 4 nights, p = 0.1; HMC scaled versus HMC, RAI -4.1 [-7.3 to -0.9], 6 nights, p = 0.03; scaled HMC in MGH's format versus scaled HMC, SINBAR -5.4 [-9.1 to -1.8], 8 nights, p = 0.008; MGH with noise versus MGH (277 nights), Montreal -23.5 [-28.5 to -18.5], 65 nights all positive to negative, p = 5 × 10⁻²⁰; SINBAR -47.3 [-53.2 to -41.4], 131 nights, p = 7 × 10⁻⁴⁰; RAI -1.4 [-2.8 to 0.0], 4 nights, p = 0.1. Storage steps read from the file headers: HMC 0.024 µV (16 bits over ±800 µV); MGH 1 µV.

Supplementary Table S5. Definition of background: share flagged (%) and discrimination of CAP RBD from same-system patients (AUC). Default: 5th percentile of 1-s REM RMS (p5 REM); p5 NREM: of N2/N3. AASM tonic keeps its NREM definition throughout.

| Background | Montreal: CAP RBD / healthy / NFLE / PLM / MGH / HMC | SINBAR: CAP RBD / healthy / NFLE / PLM / MGH / HMC | AUC Montreal phasic / tonic | AUC SINBAR phasic / any |
| --- | --- | --- | --- | --- |
| Minimum of REM | 91 / 33 / 40 / 50 / 47 / 15 | 95 / 73 / 90 / 100 / 91 / 39 | 0.869 / 0.794 | 0.799 / 0.817 |
| p5 REM (default) | 91 / 13 / 22 / 40 / 29 / 9 | 95 / 60 / 75 / 70 / 76 / 25 | 0.846 / 0.844 | 0.884 / 0.850 |
| p10 REM | 91 / 13 / 18 / 30 / 26 / 9 | 95 / 60 / 72 / 70 / 72 / 21 | 0.800 / 0.818 | 0.874 / 0.848 |
| p25 REM | 55 / 13 / 15 / 20 / 21 / 9 | 91 / 53 / 65 / 70 / 63 / 18 | 0.728 / 0.711 | 0.851 / 0.825 |
| p50 REM | 27 / 0 / 5 / 10 / 13 / 7 | 86 / 33 / 45 / 40 / 45 / 12 | 0.556 / 0.473 | 0.750 / 0.730 |
| p5 NREM | 64 / 13 / 15 / 30 / 17 / 8 | 95 / 33 / 57 / 70 / 55 / 16 | 0.751 / 0.734 | 0.843 / 0.775 |

MGH stayed above HMC under every definition. Low references separated RBD best; references near the median moved into RBD's own excess activity and separation collapsed (Supplementary Figure S2).

Supplementary Table S6. Cut-offs re-derived for automated scoring (Youden's index, CAP RBD n = 21 vs same-system patients n = 51). Cut-off with 95% bootstrap interval (2,000 re-derivations); leave-one-out (LOO) sensitivity and false-positive rate on CAP; share of MGH and HMC nights flagged [bootstrap interval].

| Measure (published cut-off) | Re-derived cut-off | LOO sensitivity, % | LOO false positives, % | MGH flagged, % | HMC flagged, % |
| --- | --- | --- | --- | --- | --- |
| Montreal phasic % (15) | 27.7 [19.7-29.3] | 86 | 18 | 8 [7-17] | 5 [5-7] |
| Montreal tonic % (30) | 32.2 [32.2-46.6] | 90 | 14 | 13 [6-13] | 7 [5-7] |
| SINBAR any % (18.2) | 62.9 [62.9-80.4] | 86 | 20 | 10 [3-10] | 1 [1-1] |
| SINBAR phasic % | 60.1 [56.9-73.4] | 86 | 18 | 9 [3-11] | 2 [1-2] |
| SINBAR tonic % | 30.9 [30.9-47.0] | 90 | 14 | 13 [6-13] | 7 [5-7] |
| AASM phasic-epoch % (none) | 31.6 [18.9-34.6] | 81 | 12 | 7 [6-18] | 4 [4-6] |
| AASM phasic % | 23.5 [23.5-34.2] | 86 | 12 | 16 [6-16] | 6 [4-6] |
| AASM tonic % | 42.0 [37.4-90.7] | 81 | 59 | 46 [9-51] | 34 [9-36] |
| RAI (0.8) | 0.92 [0.84-0.92] | 95 | 24 | 25 [12-25] | 14 [9-14] |
| SINBAR-FC any % | 49.1 | 86 | 25 | 13 [10-17] | 14 [11-20] |
| Montreal-FC phasic % | 7.9 | 86 | 27 | 32 [10-32] | 22 [10-22] |
| Montreal-FC tonic % | 6.6 | 76 | 35 | 33 [8-33] | 20 [7-20] |

Only the SINBAR-FC cut-off met the pre-specified definition of portable (difference within 5 points with LOO sensitivity at least 80%; Supplementary Figure S3).

Within-cohort quiet-REM level. Share flagged by tertile of the night's own level (tertile medians): MGH Montreal 27% (0.43 µV), 27% (0.59), 34% (0.92); SINBAR 74, 76, 78%; RAI 2, 6, 19%. HMC Montreal 0% (1.14 µV), 12% (1.52), 14% (1.96); SINBAR 18, 29, 29%; RAI 0, 6, 16%. In MGH, odds ratio per halving of the level, adjusted for age, sex and AHI: Montreal 0.64 [0.55-0.75], RAI 0.28 [0.21-0.37], SINBAR 1.20 [1.03-1.39] (Supplementary Figure S4). Within a cohort the level also carries genuine muscle tone (CAP RBD 0.80 µV against 0.47-0.60 µV in CAP's other patients).

MGH by record batch. Across the 12 undocumented record batches (n = 19-151), the batch's median quiet-REM level correlated with its SINBAR positivity at Spearman ρ -0.52 (Montreal +0.35). The first 300 records used for the ladder were quieter than the rest (0.56 vs 0.61 µV; SINBAR 83% vs 73%).

Age. Same-system patients aged 50 or over (n = 16; 59 ± 6 years) were flagged more often than younger ones (n = 35; 30 ± 9): Montreal 44% vs 29%, SINBAR 75% vs 74%, RAI 19% vs 11%, SINBAR-FC 62% vs 37%. Discrimination of RBD (n = 21, 71 ± 7 years) held against patients aged 50 or over (AUC: Montreal tonic 0.866, phasic 0.812; SINBAR any 0.884; RAI 0.814).

Heartbeat artifact. Heartbeat-locked chin amplitude ratio, median [90th percentile], before and after template subtraction: CAP 1.34 [2.65] to 1.03 [1.17]; MGH 1.92 [5.13] to 1.08 [1.53]; HMC 1.26 [2.17] to 1.01 [1.07]. By tertile of residual artifact in MGH (ratio 1.01, 1.08, 1.38): Montreal 40, 28, 21%; RAI 17, 7, 3%; SINBAR 75, 70, 84%.

Reproduction of published values. Mean ± SD (median), % of mini-epochs or epochs: SINBAR "any" 73.2 ± 21.3 (80.9) in CAP RBD against 66.5 ± 25.2 in the SINBAR development sample [5]; SINBAR tonic 51.4 ± 23.5 against 51.8 ± 32.8; SINBAR phasic 69.0 ± 20.2 against 40.9 ± 19.0. In patients without RBD, automated values were three to five times those of the published visually scored controls (SINBAR "any" 37.9 ± 24.6 in CAP's same-system patients against 8.4 ± 4.6), whose controls were treated sleep-apnea patients scored with the full SINBAR montage.

Transitional REM epochs. Without the first and last epoch of every REM period, the MGH-HMC gap went from +20 to +16 points (Montreal), +51 to +52 (SINBAR) and +1 to +1 (RAI).

Montreal phasic burst length. Limiting bursts to 10 s instead of 15 s (sources differ) changed 0 CAP RBD nights, 1 same-system patient night, 9 MGH nights and 0 HMC nights.

## S3. Noise ladder (pre-specified, commit 6d3cc6b, 26 September 2026)

Plan. Every recording was streamed again and its chin EMG kept as recorded and with white Gaussian noise added in band, after the same 10-100 Hz band-pass and mains notch as the signal. One noise track per night, seeded by record name, was scaled to each level (common random numbers): 0.5, 1.0, 1.5, 2.0 and 3.0 µV RMS. Two harmonized copies used, per night, the noise that raised the quiet-REM level to 1.5 or 2.0 µV; nights already above the target were left as they were. Cohorts: CAP (every recording with a chin EMG in physical units), HMC (all) and MGH (the first 300 records of the Challenge 2018 list, the same 300 as the earlier noise test). An independent measure, the 40-90 Hz floor of a central EEG channel in N2/N3 and REM, was computed in the same pass. The six criteria in Table 4 of the main text were committed with this plan before any ladder result existed.

Supplementary Table S7. Share flagged (%) at each ladder step, with the group's median quiet-REM level (µV).

| Group, rule | As recorded | +0.5 µV | +1.0 µV | +1.5 µV | +2.0 µV | +3.0 µV | Harmonized 1.5 µV | Harmonized 2.0 µV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CAP RBD (21): level | 0.86 | 0.99 | 1.31 | 1.71 | 2.13 | 3.02 | 1.50 | 2.00 |
| Montreal | 90 | 86 | 71 | 57 | 52 | 19 | 57 | 57 |
| SINBAR | 95 | 95 | 95 | 95 | 81 | 67 | 95 | 86 |
| RAI | 57 | 57 | 57 | 57 | 52 | 43 | 57 | 52 |
| CAP patients (51): level | 0.48 | 0.68 | 1.06 | 1.49 | 1.93 | 2.82 | 1.50 | 2.00 |
| Montreal | 33 | 18 | 16 | 12 | 12 | 10 | 14 | 12 |
| SINBAR | 75 | 53 | 37 | 24 | 22 | 20 | 25 | 24 |
| RAI | 14 | 14 | 12 | 12 | 12 | 12 | 14 | 14 |
| MGH (280): level | 0.56 | 0.74 | 1.10 | 1.52 | 1.95 | 2.84 | 1.50 | 2.00 |
| Montreal | 34 | 23 | 15 | 9 | 8 | 6 | 12 | 9 |
| SINBAR | 83 | 68 | 44 | 29 | 22 | 13 | 33 | 22 |
| RAI | 12 | 11 | 10 | 10 | 10 | 8 | 11 | 10 |
| HMC (147): level | 1.52 | 1.59 | 1.79 | 2.07 | 2.40 | 3.16 | 1.52 | 2.00 |
| Montreal | 9 | 9 | 9 | 8 | 7 | 7 | 9 | 8 |
| SINBAR | 25 | 21 | 17 | 15 | 12 | 7 | 22 | 17 |
| RAI | 7 | 7 | 7 | 7 | 6 | 6 | 7 | 7 |

Criteria in detail. (1) Harmonized to 2.0 µV, the Montreal criteria flagged 9% [6-13] of MGH and 8% [5-14] of HMC nights (+0.8 points; +25.1 as recorded) and SINBAR 22% [18-28] and 17% [12-24] (+5.5; +57.9): met for Montreal, narrowly missed for SINBAR, with 90% of the gap removed. (2) MGH's ladder interpolated at HMC's median level (1.52 µV): Montreal 8.9% against HMC 8.8%, SINBAR 29.2% against 25.2%, RAI 10.4% against 7.5%: met. (3) RBD sensitivity at a 2.0-µV floor: Montreal 57%, SINBAR 86%, RAI 52%: supported for the Montreal criteria. (4) Harmonize to 2.0 µV and re-derive cut-offs on CAP: Montreal phasic cut-off 5.4% (LOO sensitivity 76%; MGH 21% [8-47] vs HMC 10% [5-16]); Montreal tonic 1.5% (71%; 40% vs 19%); SINBAR any 13.7% (90%; 30% vs 20%): not met; harmonizing then re-deriving is not a working remedy, because winning back sensitivity drives the cut-offs into the part of the distributions where the sites still differ. (5) EEG 40-90 Hz floor in N2/N3: CAP 0.16 µV [0.15-0.20], MGH 0.23 [0.22-0.25], HMC 0.46 [0.41-0.82]; HMC/MGH ratio 1.98: met, the noise difference is system-wide. MGH's EEG floor lies near the quantization noise of its 1-µV steps in that band (about 0.2 µV), so the ratio is a lower bound. (6) Without transitional REM epochs the Montreal gap changed from +20 to +16 points: narrowly not met; REM boundaries carry about 4 of the 20 Montreal points and none of SINBAR's.

Without harmonization, the cut-offs re-derived on CAP gave, on the 300 ladder nights: Montreal phasic 27.7%, MGH 9% vs HMC 5%; Montreal tonic 32.2%, 16% vs 7%; SINBAR any 62.9%, 16% vs 1%. Harmonizing to 2.0 µV alone made MGH and HMC agree under the published cut-offs (Montreal 9% vs 8%) but cost RBD sensitivity (Montreal 90% to 57%; Supplementary Figure S5).

## S4. Floor-corrected criteria (pre-specified, commit 20b9ec2, 26 September 2026)

Derivation. A relative rule counts activity above k b. When b is noise of RMS n, a muscle signal of RMS a adds to it in power, and the rule is met when √(a² + n²) exceeds kb, that is, for a of at least about √(k² − 1)·b. The muscle activity demanded therefore scales with the system's noise floor: 1.7 b for k = 2 and 3.9 b for k = 4, about 1.0 µV at MGH's floor and 2.6 µV at HMC's for SINBAR. The floor-corrected threshold asks instead for the muscle activity the published rule asks for at a reference background B, whatever the recording's own noise. For band-limited white noise, the 5th percentile of 1-s RMS values is 0.910-0.913 times the RMS at 200, 256 and 512 Hz (simulation), hence n = b/0.91. The guard g = 1.5 was chosen on simulated noise: runs of at least 0.1 s above 1.5 b occurred in 0.4-0.7% of 3-s mini-epochs and never made an epoch tonic (at 1.3 b, 11%). The rule was developed on simulated white noise and checked on three CAP nights only before the plan was committed; the RAI, which subtracts its own noise estimate, was the comparator.

Supplementary Table S8. Floor-corrected rules: the same nights as the floor rises, and the reference background B. Share flagged (%), as recorded and at each ladder step.

| Group, rule | As recorded | +0.5 | +1.0 | +1.5 | +2.0 | +3.0 µV | Harmonized 1.5 µV | Harmonized 2.0 µV |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| CAP RBD, Montreal published | 90.5 | 85.7 | 71.4 | 57.1 | 52.4 | 19.0 | 57.1 | 57.1 |
| CAP RBD, Montreal-FC | 81.0 | 81.0 | 81.0 | 81.0 | 81.0 | 81.0 | 81.0 | 81.0 |
| CAP RBD, SINBAR published | 95.2 | 95.2 | 95.2 | 95.2 | 81.0 | 66.7 | 95.2 | 85.7 |
| CAP RBD, SINBAR-FC | 95.2 | 95.2 | 95.2 | 95.2 | 95.2 | 90.5 | 95.2 | 95.2 |
| CAP patients, Montreal published | 33.3 | 17.6 | 15.7 | 11.8 | 11.8 | 9.8 | 13.7 | 11.8 |
| CAP patients, Montreal-FC | 21.6 | 21.6 | 21.6 | 21.6 | 21.6 | 21.6 | 21.6 | 21.6 |
| CAP patients, SINBAR published | 74.5 | 52.9 | 37.3 | 23.5 | 21.6 | 19.6 | 25.5 | 23.5 |
| CAP patients, SINBAR-FC | 45.1 | 43.1 | 43.1 | 43.1 | 39.2 | 25.5 | 45.1 | 37.3 |
| MGH (280), Montreal published | 33.9 | 23.2 | 14.6 | 8.9 | 7.9 | 6.1 | 11.8 | 8.9 |
| MGH (280), Montreal-FC | 21.3 | 21.4 | 21.4 | 21.4 | 21.4 | 20.4 | 21.4 | 21.4 |
| MGH (280), SINBAR published | 83.0 | 68.2 | 43.9 | 29.3 | 22.1 | 13.2 | 33.2 | 22.5 |
| MGH (280), SINBAR-FC | 48.7 | 48.9 | 49.6 | 49.6 | 44.3 | 28.6 | 50.4 | 43.9 |
| HMC (147), Montreal published | 8.8 | 8.8 | 8.8 | 8.2 | 7.5 | 6.8 | 8.8 | 8.2 |
| HMC (147), Montreal-FC | 16.3 | 16.3 | 16.3 | 15.6 | 15.6 | 15.6 | 16.3 | 15.6 |
| HMC (147), SINBAR published | 25.2 | 21.1 | 17.0 | 15.0 | 12.2 | 7.5 | 21.8 | 17.0 |
| HMC (147), SINBAR-FC | 42.9 | 41.5 | 40.1 | 35.4 | 26.5 | 18.4 | 42.9 | 34.7 |

Paired, as recorded to harmonized 1.5 µV (same nights; exact McNemar): MGH Montreal published 33.9% to 11.9% (-22.0 [-26.9 to -17.1]; 61 nights to negative, none to positive; p = 9 × 10⁻¹⁹), Montreal-FC 21.3% to 21.3% (no night changed), clamp 19.1% to 11.9% (-7.2; 20 nights); MGH SINBAR published 83.0% to 33.2% (-49.8 [-55.7 to -43.9]; 138 nights; p = 6 × 10⁻⁴²), SINBAR-FC 48.7% to 50.5% (+1.8 [+0.2 to +3.4]; 5 nights to positive, none to negative; p = 0.06), clamp 45.1% to 33.2% (-11.9); CAP RBD Montreal published 90.5% to 57.1% (7 nights; p = 0.02), Montreal-FC unchanged. Above a floor of about 2 µV the guard binds and floor-corrected positivity falls: below that floor small activity cannot be told from noise.

| Rule | CAP RBD, % | CAP patients, % | MGH, % | HMC, % | MGH - HMC | Ladder change, MGH | Ladder change, RBD |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Montreal, published | 90 | 33 | 29 | 9 | +20.5 | -22.0 | -33.3 |
| Montreal-FC, B = 0.5 µV | 95 | 41 | 41 | 35 | +5.5 | -1.1 | 0 |
| Montreal-FC, B = 0.75 µV | 95 | 27 | 26 | 21 | +4.5 | -1.1 | 0 |
| Montreal-FC, B = 1.0 µV (primary) | 81 | 22 | 18 | 16 | +1.6 | 0.0 | 0 |
| Montreal-FC, B = 1.25 µV | 71 | 22 | 13 | 12 | +1.9 | -0.7 | 0 |
| Montreal-FC, B = 1.5 µV | 62 | 18 | 11 | 11 | -0.2 | -0.4 | 0 |
| Montreal-FC, B = 2.0 µV | 57 | 14 | 8 | 9 | -0.6 | 0.0 | -4.8 |
| Montreal clamp, B = 1.0 µV | 76 | 18 | 16 | 9 | +7.5 | -7.2 | -19.0 |
| SINBAR, published | 95 | 75 | 76 | 25 | +51.0 | -49.8 | 0 |
| SINBAR-FC, B = 0.5 µV | 95 | 71 | 78 | 52 | +25.2 | -23.1 | 0 |
| SINBAR-FC, B = 0.75 µV | 95 | 51 | 59 | 51 | +7.8 | -6.9 | 0 |
| SINBAR-FC, B = 1.0 µV (primary) | 95 | 45 | 45 | 43 | +2.4 | +1.8 | 0 |
| SINBAR-FC, B = 1.25 µV | 95 | 29 | 35 | 31 | +3.7 | -0.4 | 0 |
| SINBAR-FC, B = 1.5 µV | 95 | 27 | 28 | 22 | +5.8 | -0.7 | 0 |
| SINBAR-FC, B = 2.0 µV | 86 | 25 | 19 | 18 | +1.2 | -0.4 | 0 |
| SINBAR clamp, B = 1.0 µV | 95 | 35 | 40 | 25 | +15.2 | -11.9 | 0 |

Ladder change: as recorded to harmonized 1.5 µV, points. For B of at least 1.0 µV both rules were noise-invariant (changes within 2 points) with site differences within 6 points; B set the operating point. B = 0.5 µV failed for SINBAR: it asks for less muscle activity than HMC's floor can reveal, so the guard binds at HMC. With B = 0.75 µV, Montreal-FC kept 95% sensitivity with a 4.5-point gap (exploratory, not a pass).

Discrimination within CAP (AUC, RBD vs same-system patients). Montreal tonic: published 0.844, FC 0.827 (B = 0.5 to 2.0 µV: 0.864 to 0.788); Montreal phasic 0.846 and 0.847; SINBAR any 0.850 and 0.833. Within one system there is no floor difference to correct, so the gain from floor correction is portability, not separation.

Gain. Because floor-corrected thresholds are partly absolute, they change with gain: HMC scaled by 0.39 lowered Montreal-FC from 16.3% to 7.5% (13 nights; p = 0.0002) and SINBAR-FC from 42.9% to 11.6% (46 nights; p = 3 × 10⁻¹⁴), while the published rules were unchanged. Real sites did not differ in gain. Net levels (noise removed in power; median [IQR], µV): N2/N3 1.60 [0.93-2.90] in MGH, 1.70 [1.06-2.77] in HMC, 1.61 [1.07-2.46] in CAP's same-system patients and 2.91 [1.36-5.18] in CAP RBD; REM median 0.44, 0.45, 0.39 and 2.33; REM 95th percentile 2.91, 2.05, 2.25 and 11.77. MGH/HMC ratios were 0.95 (N2/N3), 0.99 (REM median) and 1.42 (REM 95th percentile), all within the pre-specified 0.67-1.5.

Limits of the rule. On CAP's healthy volunteers, recorded on several systems, floor correction turned three nights positive under the Montreal criteria, all on a 128-Hz system with floors of 2.5-3.7 µV. At 128 Hz the 5th percentile of noise RMS is 0.879 (not 0.91) of the RMS and the 1.5 b guard lets noise start 0.1-s runs in 2.3% of 3-s mini-epochs (0.5% at 200 Hz; 8.8% at 100 Hz), and these floors exceed the rule's range (b at most B/0.59 = 1.7 µV for B = 1.0 µV). Floor-corrected scores should be read only at sampling rates of at least 200 Hz, the AASM minimum, and within that range. Three CAP healthy recordings are stored without physical units and were not scored by floor-corrected or absolute rules.

## S5. Recording systems not used so far (pre-specified, commit 48053f3, 27 September 2026)

Plan. The plan was committed before any chin EMG of these cohorts was processed; only metadata had been read (licenses, channel lists, sampling rates, units, stage coding and the structure of one DOD-H file). Rules, code paths and parameters were those already fixed. Inclusion was unchanged: chin EMG in physical units sampled at 200 Hz or more, expert or consensus stages in 30-s epochs, and at least 10 min of REM sleep. BOAS (OpenNeuro ds005555, CC0; Micromed SD plus; bipolar chin EMG at 256 Hz; consensus AASM stages of three scorers): 128 nights of 100 adults; each participant's first recording was primary (99 with at least 10 min of REM), all nights a sensitivity analysis. DOD-O (Dreem Open Datasets, Zenodo 15900394, MIT license; Somnomedics Somno HD; 250 Hz): 56 records, one without consensus scoring and two with less than 10 min of REM excluded, leaving 53. DOD-H (same release; Compumedics Siesta; 250 Hz): 25 healthy volunteers aged 18-65. Both Dreem sets provide ECG and consensus stages of five scorers. Mains: 60 Hz for DOD-O, 50 Hz for DOD-H and BOAS. BOAS has no ECG channel: R-peaks were sought on the chin EMG itself with the same detector and plausibility test (35-130 beats/min, more than 90% regular intervals), and the heartbeat was removed only where detection passed (6 of 99 nights).

Heartbeat check (exploratory, added after the first look at BOAS). Periodicity of the REM envelope at the heart rate (spectral peak at 0.7-2 Hz over the 0.5-3 Hz median): MGH 2.67 after removal (90th percentile 4.66) and 3.17 with the heartbeat left in; HMC 2.33 and 2.50; DOD-O 2.25 and 2.74; DOD-H 2.67 and 2.61; BOAS 3.07 (90th percentile 5.61). In the 80 BOAS nights no more periodic than MGH's after removal, positivity was the same as in all BOAS nights (Montreal 44%, Montreal-FC 14%, SINBAR 92%, SINBAR-FC 35%). Leaving the heartbeat in, in the cohorts with ECG, changed the published rules by at most 7 points (it raises the background they are multiples of) but raised SINBAR-FC by 9-15 points (MGH 45% to 60%, HMC 43% to 52%, DOD-O 51% to 62%, DOD-H 32% to 40%).

Results. Share flagged and the prediction verdicts are in main Tables 3 and 4. BOAS, all 127 nights: Montreal 41% (FC 12%), SINBAR 91% (FC 35%). BOAS first nights by age: below 50 years (n = 74) Montreal 36% (FC 15%), SINBAR 89% (FC 31%); 50 or over (n = 24) Montreal 58% (FC 17%), SINBAR 96% (FC 50%). Net levels (µV): BOAS N2/N3 1.04 [0.69-2.00], REM 0.33; DOD-O 1.68 [0.91-4.04] and 0.45; DOD-H 1.05 [0.78-1.97] and 0.27. Floor signature (ratio of quietest N2/N3 to quietest REM chin level): BOAS 1.20, DOD-O 1.16, DOD-H 1.10. MGH's ladder interpolated at DOD-O's floor (0.82 µV) predicted Montreal 20.9% and SINBAR 61.5%; DOD-O was flagged at 30.2% and 58.5%.

The four narrow misses share an explanation: the two healthy cohorts, recorded on different systems, had the same lower net N2/N3 tone (1.04 and 1.05 µV, against 1.6-1.7 µV in the three patient cohorts), which points to their population rather than to gain; and the re-derived SINBAR-FC cut-off flagged 11-12% of their nights against the 10% or less set in advance, so it is about 88% specific in healthy adults.

## S6. Parkinson's disease analyses (pre-specified, commit 1e7acf0, 2 October 2026)

Plan. The plan was written and committed with its code before any RBDtector run. Isolated RBD is the strongest known prodromal marker of Parkinson's disease, dementia with Lewy bodies and multiple system atrophy [1], and RWA is used in five ways in that research: diagnosis and trial eligibility; conversion risk (baseline tonic chin RWA 72.9% in later Parkinson's disease against 41.0% in those who stayed disease-free, 62.7% for any neurodegeneration [19]); progression (30% to 54% over a mean 5 years in 11 patients [20]); an earlier stage, isolated RWA (DREAMER, RAI below 0.85 [18]; 25% RBD, 25% isolated RWA and 50% normal in early Parkinson's disease [21]); and dementia with Lewy bodies, where polysomnographic RWA is an indicative biomarker [22]. No dataset with Parkinson's patients is available without an application, so nothing here is measured in people with Parkinson's disease. The ten predictions are listed in main Table 4.

RBDtector setup. RBDtector (commit 7c9c73d, 19 May 2022) was run through its own controller (PSGController.run_rbd_detection) with default settings, on input files written in its format: the chin EMG as channel "EMG" in microvolts, band-passed at 10-100 Hz and notched as an acquisition system would (RBDtector does not filter), heartbeat left in, at 256 Hz (CAP decimated from 512 Hz, which is what RBDtector's own spline resampling does at that rate; MGH's 200 Hz resampled by RBDtector); the expert stages as its sleep profile; and no arousals (one dummy event in the first second, which is wake). It ran in the environment its requirements specify (Python 3.9, NumPy 1.21, pandas 1.3, SciPy 1.7, pyEDFlib 0.1.42, the newest release its pin allows, whose renamed sampling-rate key was mapped back to the name RBDtector reads). Two changes made the 1,412 runs fit in Colab sessions without changing any score: a helper that assigns each short REM fragment to the nearest REM block was replaced by the equivalent NumPy operation, and each night was written from two epochs before its first REM epoch to two after its last, keeping the night's clock times so that RBDtector's 30-ms, 3-s and 30-s grids fell where they would. On a test night the scores were identical, to the last digit, to those of the unchanged program on the whole night, and a run took 35 s instead of 81 s. The floor-corrected run of each night reused the baseline its published run had computed from the same files; the ratio of RBDtector's baseline to the RMS of band-limited white noise (0.976) was obtained by simulation.

Deviations from the run plan (none changed a score). All 1,412 runs were completed (72 CAP nights at five ladder steps; 100 MGH, 147 HMC and 99 BOAS nights; each with both baselines); no night failed and no ladder step was dropped. The work was split between two Colab runtimes; the second had no Google Drive access, so its record lists were passed to it from the first and its scores file was returned through the browser, with an identical MD5 checksum and exactly the pre-specified MGH and HMC nights. The CAP run stopped once after 61 of 72 nights when a second job exhausted the runtime's memory, and was resumed; scores are written only when a night is complete.

Supplementary Table S9. Per-night change from the night as recorded to a quiet-REM level of 1.5 µV, percentage points (mean [95% bootstrap CI]).

| Measure | CAP RBD (21): as recorded to 1.5 µV | Change | CAP patients (51): as recorded to 1.5 µV | Change |
| --- | --- | --- | --- | --- |
| Montreal tonic, published | 53.3 to 31.8 | -21.4 [-29.7 to -13.2] | 17.9 to 9.8 | -8.1 [-10.4 to -6.0] |
| Montreal phasic, published | 37.1 to 22.8 | -14.3 [-18.4 to -9.7] | 13.8 to 7.8 | -6.0 [-7.7 to -4.5] |
| SINBAR any, published | 74.8 to 54.1 | -20.7 [-28.9 to -12.4] | 37.9 to 20.2 | -17.7 [-21.7 to -14.1] |
| RBDtector tonic, published | 40.3 to 28.3 | -12.1 [-17.0 to -7.3] | 13.2 to 6.9 | -6.3 [-8.2 to -4.5] |
| RBDtector phasic, published | 31.4 to 27.3 | -4.1 [-8.1 to -0.6] | 23.8 to 12.9 | -10.9 [-14.0 to -8.4] |
| RBDtector any, published | 69.0 to 54.9 | -14.1 [-19.4 to -8.7] | 36.9 to 20.1 | -16.8 [-20.4 to -13.5] |
| Montreal tonic, FC | 44.7 to 43.6 | -1.1 [-1.7 to -0.5] | 14.0 to 13.8 | -0.2 [-0.4 to 0.0] |
| Montreal phasic, FC | 39.4 to 39.0 | -0.4 [-0.6 to -0.2] | 12.1 to 12.3 | +0.2 [-0.1 to +0.6] |
| SINBAR any, FC | 68.7 to 67.5 | -1.1 [-1.9 to -0.5] | 27.3 to 27.4 | 0.0 [-0.2 to +0.3] |
| RBDtector tonic, FC | 47.4 to 47.5 | +0.1 [-0.9 to +1.3] | 12.1 to 12.7 | +0.6 [+0.2 to +0.9] |
| RBDtector phasic, FC | 33.7 to 33.2 | -0.4 [-1.2 to +0.3] | 19.2 to 21.5 | +2.3 [+1.2 to +3.3] |
| RBDtector any, FC | 77.0 to 76.9 | -0.1 [-1.0 to +0.6] | 31.1 to 34.0 | +2.9 [+1.7 to +4.1] |

Against the benchmarks, the published Montreal tonic change in RBD equals 67% of the 31.9-point difference that predicted Parkinson's disease and 4.5 years of progression at 4.8 points per year; RBDtector's equals 38% and 2.5 years; the floor-corrected changes equal 3% and 0%. Mean tonic % of CAP RBD along the ladder (as recorded, +0.5, +1.0, +1.5, +2.0, +3.0 µV): Montreal published 53.3, 45.9, 33.9, 25.6, 19.8, 11.3; Montreal-FC 44.7, 44.4, 43.3, 41.5, 35.2, 25.0; RBDtector published at as recorded, +0.5, +1.0 and +2.0 µV 40.3, 37.5, 31.0, 20.1; RBDtector-FC 47.4, 48.0, 48.5, 38.7. Beyond a floor of about 1.7 µV the floor-corrected rules leave their valid range.

RBDtector along the ladder, % flagged. CAP RBD (quiet-REM level 0.86, 0.99, 1.31, 2.13 µV at as recorded, +0.5, +1.0, +2.0 µV; 1.50 µV harmonized): published 95.2, 95.2, 95.2, 85.7 and 95.2; FC 95.2 at every step (no night changed). CAP patients (0.48, 0.68, 1.06, 1.93; 1.50 µV): published 78.4, 47.1, 31.4, 19.6 and 25.5 (-52.9 points [-66.6 to -39.2] as recorded to 1.5 µV; 27 nights to negative, none to positive; p = 1 × 10⁻⁸); FC 37.3, 39.2, 45.1, 37.3 and 49.0 (+11.8 [+2.9 to +20.6]; none to negative, 6 to positive; p = 0.03). Agreement with our SINBAR "any" (published), same 418 nights as recorded: Spearman ρ +0.79; RBDtector's baseline was 1.14 times our quiet-REM floor (median).

Supplementary Table S10. Isolated RWA: share of nights without RBD labeled by each definition (%).

| Group | Nights | Floor, µV | RAI below 0.85 (DREAMER) | RAI below 0.8 | SINBAR published / FC | Montreal published / FC | RBDtector published / FC |
| --- | --- | --- | --- | --- | --- | --- | --- |
| CAP same-system patients | 51 | 0.48 | 16 | 14 | 75 / 45 | 33 / 22 | 78 / 37 |
| CAP patients aged 50 or over | 16 | 0.53 | 25 | 19 | 75 / 62 | 44 / 25 | 75 / 62 |
| MGH | 925 | 0.59 | 12 | 9 | 76 / 45 | 29 / 18 | 65 / 55 |
| MGH aged 50 or over | 610 | 0.59 | 12 | 8 | 75 / 43 | 30 / 17 | 69 / 53 |
| HMC | 147 | 1.52 | 9 | 7 | 25 / 43 | 9 / 16 | 20 / 56 |
| DOD-O | 53 | 0.82 | 13 | 8 | 58 / 51 | 30 / 25 | - |
| BOAS, healthy | 99 | 0.37 | 15 | 13 | 91 / 35 | 41 / 15 | 89 / 34 |
| BOAS aged 50 or over | 24 | 0.38 | 17 | 17 | 96 / 50 | 58 / 17 | 96 / 54 |
| DOD-H, healthy | 25 | 0.56 | 8 | 4 | 84 / 32 | 20 / 12 | - |

RAI below 0.85 on the noise ladder (same nights; as recorded, at 1.5 µV, and at +3.0 µV): CAP patients 16, 16 and 14%; MGH 15, 13 and 11%; HMC 9, 9 and 8%; CAP RBD 76, 67 and 62%. RBDtector columns: MGH's random 100 nights.

Supplementary Table S11. (a) Recruitment for prevention trials per 1,000 PSGs of people aged 50 or over at an RBD prevalence of 2%. False-positive rate from MGH patients aged 50 or over (an upper bound; RBDtector: 68 of its random 100), sensitivity from CAP RBD (n = 21); expected to develop a synucleinopathy within 12 years: 73.5% of those found [1].

| Rule | False positives, % | Sensitivity, % | Flagged | RBD found | Positive predictive value, % | Studies per RBD found | Expected to convert |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Montreal, published | 29.5 | 90 | 308 | 18.1 | 5.9 | 17 | 13.3 |
| Montreal-FC | 17.3 | 81 | 186 | 16.2 | 8.7 | 11 | 11.9 |
| SINBAR, published | 74.8 | 95 | 752 | 19.0 | 2.5 | 39 | 14.0 |
| SINBAR-FC | 42.9 | 95 | 440 | 19.0 | 4.3 | 23 | 14.0 |
| SINBAR-FC at least 49.1% | 11.4 | 86 | 129 | 17.1 | 13.3 | 8 | 12.6 |
| RAI below 0.8 | 8.4 | 57 | 93 | 11.4 | 12.2 | 8 | 8.4 |
| RBDtector, published | 69.1 | 95 | 696 | 19.0 | 2.7 | 37 | 14.0 |
| RBDtector-FC | 52.9 | 95 | 538 | 19.0 | 3.5 | 28 | 14.0 |

Across 1-5% prevalence, the positive predictive value ranged from 1.3-6.3% (SINBAR, published) to 7.1-28.4% (SINBAR-FC at least 49.1%). With HMC's false-positive rate instead (all ages, 2%), the published Montreal criteria would flag 105 and find 18.1 (17.3%), but that uses CAP's sensitivity, which HMC's floor lowers (57% on the ladder; then about 11 found and a positive predictive value of about 12%).

(b) Floor check by recording system. Floor b: 5th percentile of 1-s REM RMS. In range: sampling rate at least 200 Hz and b at most 1.69 µV. Muscle needed: RMS of muscle activity, on top of the noise, that a threshold of 2b demands; the floor-corrected rule demands 1.73 µV on every system.

| Group | Nights | Sampling rate, Hz | Floor, µV [IQR] | In range, % | Net N2/N3, µV | Muscle needed at 2b, µV |
| --- | --- | --- | --- | --- | --- | --- |
| CAP RBD | 21 | 512 | 0.86 [0.61-1.24] | 76 | 2.91 | 1.44 |
| CAP same-system patients | 51 | 512 | 0.48 [0.44-0.69] | 86 | 1.61 | 0.81 |
| MGH | 925 | 200 | 0.59 [0.47-0.77] | 95 | 1.60 | 0.99 |
| HMC | 147 | 256 | 1.52 [1.23-1.79] | 63 | 1.70 | 2.54 |
| DOD-O | 53 | 250 | 0.82 [0.76-0.92] | 92 | 1.68 | 1.37 |
| BOAS | 99 | 256 | 0.37 [0.28-0.49] | 92 | 1.04 | 0.61 |
| DOD-H | 25 | 250 | 0.56 [0.54-0.65] | 100 | 1.05 | 0.94 |

The same check is provided as a command-line tool (check_floor.py) that takes one EDF recording and its 30-s stages.

Exploratory: RBDtector on noise alone (added after PD2b was missed). Twelve simulated nights of band-limited white noise and nothing else (10-100 Hz with a 50-Hz notch, 256 Hz; 80 min of REM in four 20-min periods), three per noise level. Share of REM mini-epochs counted as "any" activity, mean of three nights (highest in brackets); tonic activity was 0 everywhere.

| Noise RMS, µV | 0.5 | 1.0 | 1.5 | 2.0 |
| --- | --- | --- | --- | --- |
| RBDtector baseline, µV | 0.46 | 0.95 | 1.43 | 1.92 |
| Published baseline | 0 | 0 | 0 | 0 |
| Floor-corrected, guard 1.5 (pre-specified) | 0 | 0 | 0.40 (0.62) | 1.06 (1.62) |
| Floor-corrected, guard 2.0 | 0 | 0 | 0 | 0 |

Noise alone accounts for about 0.4 of the 2.9-point rise in "any" among CAP's patients at 1.5 µV. The rest presumably comes from small muscle activity just below the corrected threshold, which added noise lifts over it more often than it pushes larger activity below it; this was not tested further. A guard of 2.0 removed the little activity noise produced alone, but its effect on real nights was not tested, and at floors near 1.5 µV it would raise the threshold above what the floor-corrected rule is meant to demand.

## S7. STARD 2015 checklist and 2 x 2 tables

Supplementary Table S12. (a) Cross-tabulation of index tests against the reference standard on CAP's main system. Reference standard: clinical diagnosis of RBD in CAP. RBD, n = 21; same-system patients without RBD, n = 51; nights as recorded.

| Index test | RBD: positive | RBD: negative | No RBD: positive | No RBD: negative | Sensitivity, % | Specificity, % |
| --- | --- | --- | --- | --- | --- | --- |
| Montreal, published | 19 | 2 | 17 | 34 | 90.5 | 66.7 |
| SINBAR, published | 20 | 1 | 38 | 13 | 95.2 | 25.5 |
| RAI below 0.8 | 12 | 9 | 7 | 44 | 57.1 | 86.3 |
| Montreal-FC | 17 | 4 | 11 | 40 | 81.0 | 78.4 |
| SINBAR-FC | 20 | 1 | 23 | 28 | 95.2 | 54.9 |
| SINBAR-FC at least 49.1%* | 18 | 3 | 10 | 41 | 85.7 | 80.4 |
| RBDtector, published | 20 | 1 | 40 | 11 | 95.2 | 21.6 |
| RBDtector-FC | 20 | 1 | 19 | 32 | 95.2 | 62.7 |

*Cut-off derived on these nights, so these values are optimistic; leave-one-out estimates are 86% sensitivity and 75% specificity (Supplementary Table S6).

(b) STARD 2015 checklist.

| Item | Section | Where reported |
| --- | --- | --- |
| 1 | Title or abstract | Title names automated RWA scoring; abstract reports sensitivity and false-positive rates |
| 2 | Structured abstract | Abstract |
| 3 | Background, intended use and clinical role | Introduction (screening routine PSGs for prodromal synucleinopathy; multi-center research) |
| 4 | Objectives and hypotheses | Introduction, last paragraph; four pre-specified plans (Supplementary S3-S6) |
| 5 | Prospective or retrospective | Methods: retrospective analysis of pre-existing open databases |
| 6 | Eligibility criteria | Methods: chin EMG present; at least 10 min of REM; physical units for absolute and floor-corrected rules |
| 7 | Basis for identifying participants | Methods: all records of each database |
| 8 | Setting, location, dates | Table 1 (places); recording periods are not stated by the databases |
| 9 | Consecutive, random or convenience series | CAP: convenience series as released; MGH and HMC: as released; RBDtector on a random 100 MGH nights (seed 0) |
| 10a | Index test, in replicable detail | Methods, Table 2, public code |
| 10b | Reference standard, in replicable detail | Methods: clinical diagnosis of RBD in CAP; the database does not state its criteria |
| 11 | Rationale for the reference standard | Methods: CAP holds the only openly available recordings of patients with RBD |
| 12a | Index test cut-offs: pre-specified or exploratory | Published cut-offs pre-specified; re-derived cut-offs exploratory; floor-corrected rules pre-specified (commit 20b9ec2) |
| 12b | Reference standard cut-offs | Not applicable (clinical diagnosis) |
| 13a | Index test blinded to reference standard | Yes: fully automated scoring |
| 13b | Reference standard blinded to index test | Unknown for CAP; the diagnosis itself requires polysomnographic RWA (incorporation bias, Discussion) |
| 14 | Methods for estimating accuracy | Methods: Statistical analysis |
| 15 | Indeterminate results | None; three CAP healthy recordings without physical units were not scored by absolute or floor-corrected rules (Supplementary S4) |
| 16 | Missing data | No chin EMG in 1 record; HMC demographics only as group summaries; no ECG in BOAS (heartbeat removed where detectable on the EMG) |
| 17 | Analyses of variability | Pre-specified analyses labeled with their commits; all others labeled exploratory |
| 18 | Intended sample size | All available open data; precision reported with 95% intervals |
| 19 | Flow of participants | Supplementary Figure S1 |
| 20 | Baseline characteristics | Table 1 |
| 21a | Severity of the target condition | Established clinical RBD (spectrum bias, Discussion); RWA distributions in Supplementary S2 |
| 21b | Alternative diagnoses | Table 1 (CAP diagnostic groups); MGH and HMC: clinical populations, diagnoses not released |
| 22 | Interval between index test and reference standard | Index tests from the diagnostic recording; interval to diagnosis not stated by CAP |
| 23 | Cross-tabulation | Supplementary Table S12a |
| 24 | Estimates of accuracy and precision | Table 3; Supplementary Tables S1, S6 and S12a |
| 25 | Adverse events | Not applicable (secondary analysis of existing recordings) |
| 26 | Limitations | Discussion: Strengths and limitations |
| 27 | Implications for practice | Discussion; Conclusions |
| 28 | Registration | Not registered in an external registry; four analysis plans committed to the project repository before they were run (Declarations) |
| 29 | Protocol | Repository: the four analysis plans and their code |
| 30 | Funding | Declarations |

## Supplementary figure legends

Figure files (Google Drive folder rbd-rwa/figures): Figure S2 = fig3_background_definition, Figure S3 = fig4_cutoffs, Figure S4 = fig5_level_vs_verdict, Figure S5 = fig7_harmonisation. Figure S1 is the flow diagram in Section S1.

Supplementary Figure S1. Participant flow. Recordings in each database, exclusions and nights analyzed (Section S1).

Supplementary Figure S2. Sensitivity of the verdicts to the operational definition of background. Every night re-scored with the background of the relative criteria set to the minimum or the 5th (default), 10th, 25th or 50th percentile of 1-s REM RMS amplitude, or the 5th percentile of 1-s N2/N3 amplitude. (A, B) Share of nights flagged positive by the Montreal and SINBAR criteria, by group. (C) Separation of CAP RBD from CAP patients recorded on the same system (AUC) for the continuous tonic and phasic measures; the AASM tonic criterion keeps its NREM reference throughout and is not shown.

Supplementary Figure S3. Published and re-derived cut-offs against each group's distribution. (A) Montreal phasic activity (% of 2-s mini-epochs; the phasic criterion alone) and (B) SINBAR "any" activity (% of 3-s mini-epochs), one point per night. Dashed line: published cut-off; solid line: cut-off re-derived for automated scoring by maximizing Youden's index for CAP RBD against CAP patients recorded on the same system. Beneath each group: share of nights flagged by the published and by the re-derived cut-off.

Supplementary Figure S4. Positive calls by a night's own quiet-REM level. Share of nights flagged positive by tertile of the quiet-REM level within (A) MGH and (B) HMC. A lower level did not make a night more likely to be flagged by the relative criteria; the RAI, built on absolute amplitude, flagged nights with a higher level more often. Within a cohort the level also carries genuine muscle tone.

Supplementary Figure S5. Harmonizing the noise floor is not a remedy. (A) Published cut-offs, as recorded and with every night harmonized to a quiet-REM level of 2.0 µV: harmonization brings MGH and HMC together but lowers RBD sensitivity (Montreal 90% to 57%). (B) Cut-offs re-derived on harmonized CAP data (Youden's index, RBD vs same-system patients) and applied to harmonized MGH and HMC: share of external nights flagged, with 95% intervals from 2,000 bootstrap re-derivations, and leave-one-out sensitivity on CAP. The low re-derived thresholds reopen the difference between the hospitals.